EOD Machine General Instructions for Use

(Anya Goldina, 2008)

III. Calibrating tanks

1) Why is it necessary to calibrate tanks?

To ensure that all tanks record EODs accurately and consistently, tanks need to be calibrated on a regular basis. Because recorded amplitude and frequency are sensitive to changes in water temperature and conductivity, it is imperative that you and the EOD machine can account for these changes by calibrating the tanks anytime you change recording parameters. Also, since the electric field changes depending on water levels, once a tank is calibrated at a certain water level, try to maintain the tank at that level throughout the duration of your experiment (you can do this by marking the original water levels with a piece of tape, as the water evaporates add a little at a time.)

- When is it necessary to calibrate tanks?

  - When you are starting a new experiment

  - When you change either the conductivity or water levels of an existing experiment.

  - When you put in new fish into an existing experiment (increasing sample size)

  - Anytime you change anything about the physical aspects of the tanks (i.e., electrodes, ceramic tube, mesh partitions) that would require taking some aspect of the tank apart.

- Important things to know BEFORE using the calibrator.

- The calibrator is VERY fragile. You must treat it as gently as possible! If it breaks, immediately tell Phil. If you are unsure about using it, ask someone in the lab to help you. However, accidents do happen, and if the calibrator does break it is your responsibility to fix it and make sure that it works!

- When you are calibrating tanks, no data will be collected for the period of time it takes you to calibrate.  Therefore, if you are sharing the EOD room with someone else, it is important to coordinate calibrating times.  When you finish calibrating, restart the EODrecorder12, otherwise the new information from the calibrator will not be integrated into the program and you will not be able to collect any data. Simply resuming EODrecorder12 by switching to Run after Paus[e]ing, will not work!!!

- Only calibrate tanks when the water level, temperature and conductivity meet the specifications of your experiment. There is no point at calibrating a tank at one condition and testing fish at another.

- Calibration Protocol for tanks in the EOD room

(Anya Goldina, 11 Jun ‘07, adapted from V. Salazar (4 Oct.’01) and J. Herrick (2 Jul ‘02))

All MATLAB commands are italicized

  - Select MATLAB icon on the desktop to open a window

  - Type calibrator ( the calibrator window will open

    - If you or someone else is collecting data, click Pause in the EODrecorder12 window.

    - The calibrator window should have 4 buttons (calibrator, tank, calculate and offset). If you don’t see all four, close and reopen the program.

  - Take the part of the calibrator apparatus that has a transformer (blue box) connected to a BNC cable and connect the cable to the Out-1 of the RP2-#1.  This portion of the apparatus should remain plugged into the RP2-#1, Out-1 regardless of the tank that you are trying to calibrate.

  - In the EOD room, install the calibrator apparatus inside the tank of interest. (NOTE: be very careful with the calibrator; if you break any of the fine glass tubes, Phil will kill you!!)

  - When installing the calibrator, make sure that the two metal balls are to your right and that they are aligned vertically with the ceramic tube.  Also, make sure that the calibrator itself is centered with respect to the ceramic tube.

  - Connect the blue cable on the calibrator to the Type 1(tank) cable of the tank of interest.

  - At the computer, on the ‘Calibrator’ window, select the tank # to be calibrated and the gain of the amplifier.

  - Then select calibrator. Record the given value in a notebook.

  - In the EOD room, remove the blue Type 1 cable from the apparatus and reconnect it to the tank. Leave the calibrator in the tank.

  - At the computer, in the ‘Calibrator’ window, select ‘Tank’ then ‘Calculate’.

  - Two values will come up:  the ‘calib’ value and the ‘offset’ value. Record both values in your notebook.

  - When all the tanks have been calibrated, click on ‘Edit setup.m’ in the ‘Calibrator’ window.

  - Enter in the new ‘calibs’ values for each tank that you calibrated, then enter in the new ‘offsets’ values. Make sure to enter date of calibration and person doing the calibration.

  - Go to FILE and ‘save’ and close the current file only! (the lower ()  If you close the top one, it shuts down the program!

  - Select ‘close’ on the ‘Calibrator’ window if you are done for the day.  Unplug the BNC cable from the RP2 and plug back in the original cable.

  - If you need to calibrate other tanks, go back to step c and repeat.

  - For the new calibration values to be uploaded, restart eodrecorder, by typing in eodrecorder12. If the eodrecorder has been paused while you were calibrating, close the program and restart by typing in eodrecorder12.

  - Before leaving, verify that the EOD machine is collecting data from tanks where you or your labmates have fish.

Troubleshooting:

- When I start the program calibrator, not all buttons show up

  - Restart calibrator

- Calibration values are sporadic

  - Make sure that your conductivities are within the same range between tanks.

    - If you are running different species that require different conductivities, your calibration values will also be different. Higher conductivities will give you higher calibration values. However, all values should be below 1.

  - Make sure that neither the XLR connectors nor the BNC connectors are touching each other or the copper-mesh cage around the tank. Metal-metal contact will interfere with the calibrator.  Try putting a dry cloth or towel underneath the metal part of the electrodes and calibrator.

- Calibration values are greater than 1

  - You have tank (type1) and tube (type2) cables mixed up.  Recalibrate making sure that you connect the calibrator apparatus to the actual tank cable, NOT tube. The same applies for reconnecting the actual electrodes back to the cables.

- My conductivities are right, and my cables are connected properly, but I still can’t get a good calibration value on some tanks.

  - It is possible that one of the tank electrodes is bad. A simple fix is to change the electrodes. If there are no spare electrodes, take the existing pair out and see if the wire needs to be re-soldered (refer to the electrode protocol.)

IV. USING EODRECORDER

EODrecorder12 is a program written by Philip Stoddard and Michael Markham that allows you to record EODs of pulse and wave type electric fish continuously for extended periods of time.

To start using the eodrecorder you need to have the following items:

- Calibration and offset values for each tank that you intend to use (see calibration protocol above)

- Assign names to the fish whose signals you will be measuring.

  - Names are assigned based on Genus, Species, Sex and number

    - For example

      - bpm025 stands for Brachyhypopomus pinnicaudatus male, number 25

      - gcx033 stands for Gymnotus carapo, sex unknown, number 33.

      - To find out which number to give your fish, look at the bottom of the dry-erase board, find the abbreviation for your species and choose the next number up.

- Know whether your fish is pulse or wave.

- Assign a number to your experiment. As in fish numbers, look at the dry-erase board to see the last experiment number that was used and choose the next number up.

USING EODRECORDER12 (All MATLAB commands are italicized)

- Double click MATLAB icon on the desktop to open Matlab

- In the command window (largest window) type edit setup and press Enter ( a new Editor window will open.

- CREATING FILES/ FISH FOR RECORDING

Scroll down until you see 12 lines “fish1= ‘’;…………fish 12 = ‘’;”

  - These represent the 12 tanks in the EOD room into which your fish will go (i.e. fish1 is tank 1, fish2 is tank 2…..)

  - Type in the name of your fish next to the tank number in which you will put your fish.

  - Make sure that each name has single quotes on each side of it ‘bpm001’

    - If the quotes are not there, the fish name will be black.

    - If you are missing one quote, text will appear red.

  - If no one is using the first tank, the file name of tank1 should be ‘dummy’, it should look like this: fish1 = ‘dummy’. Without this, the eodrecorder will not work.

- SETTING GAINS

Scroll further down until you reach a section that says gains (1)….gains (12). Here you will set the gain at which you need to record your fish.  The gain number should reflect the gain setting of the amplifiers of each tank from which you are recording.

  - The amplifiers are located to the left of the computer. Each amplifier has a number on it that corresponds to the tank in the EOD room whose signal it is amplifying.

  - Once you have located the correct amplifier, check the setting of the gain. (It is usually set at 500).

  - Compare that setting to the one noted in the setup file.

    - If the setting is correct, you don’t have to do anything.

    - If the values don’t agree, type in the correct value next to your tank, followed by a semicolon

      - Note, no quotes are needed here.

- RECORDING WAVE vs. PULSE TYPE FISH

Scroll further down until you reach a section that says tanktype (1) …. tanktype (12).    In this section you will tell the eodrecorder whether it is collecting data from a pulse a wave-type fish.

  - While you can record EODs of wave and pulse type fish simultaneously, tank types must be in pairs.  Tanks are paired in order, tank1 is paired with tank2, tank3 with tank4 and so on. Therefore, if in tank 1 you are recording from a pulse fish, you cannot put a wave fish into tank 2.  In such scenario, you have to put a wave fish into any of the other available tanks that are not paired with “pulse” tanks.

  - Next to the corresponding tanktype type in 1 if you have pulse fish and 2 if you recording wave fish.

    - Thus, if you recording the EOD of one wave fish in tank 1 and another pulse fish in tank 4, your command lines should look as follows:      tanktype(1) =  2;  %wave

tanktype (2) = 2;  %wave

tanktype (3) = 1;  %pulse

tanktype (4) = 1;  %pulse

    - Note, even though, I only have fish in two tanks, all four tanks are assigned a pulse or wave command.

    - Signal type designation should be followed by a semicolon.

    - Any text that is preceded by a percent sign  % , is not read by Matlab and can be used to make notes or reminders for yourself.

- SETTING CALIBRATION VALUES

See calibration protocol above

- UPLOADING NEW INFORMATION AND STARTING YOUR EXPERIMENT

  - Save all your information by choosing File (upper left-hand corner) and click Save.

  - Close the Editor window.

  - In the Matlab command window type eodrecorder12

    - A new window will open that has 12 squares on the bottom, for each tank in the EOD room.

    - The upper half of the window should have a Threshold setting, a larger window in the middle and a History window on the right.

  - If you already put your fish in the tanks, then within a few seconds an EOD waveform should appear in one of the bottom windows corresponding to the tank in which you have fish.

  - If you have not put your fish in the tanks yet, go ahead and do so now. You should see the EOD signal in one of the window.

- SETTING THRESHOLDS

In order to be able to record the EODs of your fish accurately, you need to set the thresholds as tight as possible. Setting thresholds tells the eodrecorder when the fish is centered in the tube and triggers the tank electrodes to record the EOD. If you set your thresholds too high, then you might not record any data. However, if your thresholds are too low, then your data will be noisy, because any movement of the fish will trigger the eodrecorder.  Here are a few steps to help you set your thresholds:

  - Once you put your fish in the tank, wait until it enters the tube.

  - Once you are sure your fish is in the tube, come outside and look at the oscilloscope.

    - Make sure to select the proper tank by turning the dial that says “Select” below the oscilloscope until the green light is right below the right tank number.

    - Look at the signal on the screen. The signal should be in two colors, pink and blue.

    - Note where a yellow line is in relation to the signal- this is the threshold line.

      - The threshold line should be at the tip of signal. If it is too high, then you will not be able to record anything.

      - If your line is too low, then the signal will be too noisy.

    - You can adjust the threshold within the eodrecorder12 window:

      - In the top left corner that says Thresholds, choose the tank to be adjusted.

      - Increase or decrease your day settings, by pressing + or – signs and then clicking Apply. There are two ways to tell that your new thresholds have been uploaded:

        - In the History window “ thresholds changed locally and uploaded” should appear

        - On the oscilloscope, the yellow line will move in the direction of your adjustments.

    - Once you set the daytime threshold settings, adjust the nighttime thresholds.

      - A general rule for setting nighttime thresholds is that nighttime values should be double of the daytime settings. Thresholds need to be higher because the fish move during the night and if settings are not tight enough, they will trigger the eodrecorder everytime the fish is close to the tube electrodes, and not when they are centered within the tube.

        - You might have to tweak these values to make your data as accurate as possible, but this is good beginning for the first night.

- VIEWING THE SIGNAL

You can view the data you have collected by using eodshow. This is a program for viewing how the EOD of your fish changes over a period of time. You can also see different parameters of the EOD signal (amplitude, frequency, duration, etc.)

  - In the Matlab command window, type eodshow

    - A new window should open

    - On top of the window you should see a button that says D:\data – this is where these data are accessed from.

    - To your right there are two buttons:

      - Update file list – click this when you want to view a signal of a fish you have recently started recording

      - Update plot – click this when you want to see recent changes in the EOD of your fish

      - Use the scroll down menu to choose the fish whose signal you want to see.

      - Once you choose your fish, choose the EOD parameter that you want to see.

- SAVING YOUR DATA

It is imperative that you save your data at the end of every experiment and also back it up on the lab server!!!

  - While you run the experiment your data is collected on the D drive.

    - You can find it on the “Shortcut to data” folder on the desktop.

  - For every fish that you run, eodrecorder will create 3 files:

    - For example,

      - for fish bpm666, eodrecorder will create bpm666, bpm666.ana and bpm666.raw

    - At the end of your experiment, you need to save all these files.

  - To save files:

    - pause the EODrecorder12

    - Open “Shortcut to data” file on desktop and locate your files

    - Open “T Drive (data on server)” shortcut, also on desktop

      - Open EOD Machine folder

      - Open Saved_data folder

      - Create a folder within the saved_data folder entitled “Expt###”

    - Copy your data files from “Shortcut to data” folder into “Expt###” folder (this might take a few minutes, depending on the size of your files).

    - Once copied, burn this experiment unto a CD or a DVD

      - CDs and DVDs are on the desk, above the MAC, in HLS 223

      - You can burn the data on the PC in HLS 223 using Nero

      - Label the CD/DVD with your experiment number, your initials and date

      - Place the CD/DVD in the CD case labeled Expt 105(,  in the EOD room (above the PC).

    - Erase the data from the D drive once you have verified that your data is burnt on a CD and is backed up on the server.

      - Do not erase files that don’t belong to you or don’t have the names of your fish on them!

  - Open “T Drive (data on server)” shortcut on the desktop

    - Open EOD Machine folder

    - Open Experiments file

    - Create an excel file listing all the details of your experiment

      - Calibration values of your tanks

      - Weights, lengths, names of your fish

      - Dates, times, descriptions of treatments

      - Any miscellaneous info. Pertinent to your experiment

    - Place this file in the Experiments file on the server.

- GENERAL GUIDELINES FOR RUNNING EXPERIMENTS IN THE EOD MACHINE

  - First thing in the morning and last thing at night, check your threshold settings and eodrecorder to make sure that you are collecting data.

    - Especially in the summer, because of power outages, the eodrecorder tends to shut off. Therefore, if you don’t want to lose precious data, monitor the EODrecorder as often as possible.

    - Record all changes on the board, behind the PC- this will prevent confusion if more than one person is using tanks

    - Keep very good notes of everything you do from one experiment to the next.

    - If you are running an experiment for longer than 1week, rename your files. This is necessary to prevent the EODmachine from slowing down.

      - If you are running bpm666 for the first week, rename the file bpm666a for the next week, and to bpm666b for the week after.

      - The original bpm666 will be saved on the “shortcut to data” file.

      - At the end of your experiment you will end up with multiple files, bpm666, bpm666a, bpm666b that will all be of the same fish, just different times.

- POTENTIAL PROBLEMS

If you encounter any of these contact Phil

  - EODrecorder stopped collecting data

    - Restart eodrecorder

    - Restart Matlab if the first option doesn’t work

    - Contact Phil

  - The lights in the EODshow are wrong

  - Matlab says “block size is off”

    - The RP2s need to be restarted- ask Phil

  - Error message “Problem with RP2s”

    - Same as c – get Phil

  - Thresholds don’t get uploaded when you change them

    - Possible problem with server or internet connection

    - Possible problem with RP2s